## Slide 1: CLIA Waived Testing for Physician Office Labs
Doris Thompson, Ph.D., MT (ASCP)
AK State Public Health Laboratories
State CLIA Program
Francisca Lehr, M.S., MT (ASCP)
CLIA Laboratory Surveyor and Consultant
Region X CMS

## Slide 2: Introduction
What would you like for us to cover in the presentation today?
Count: reference lab, hospital, POL., large clinic, etc.
Background: CMA, RN, LPN, etc.

## Slide 3: Objectives
Learn how to find the FDA complexity of your kits/test systems
Learn what CLIA requires of waived laboratories
Learn what following the manufacturer’s instructions includes
Learn the difference between built-in controls and external controls.
Learn to apply good laboratory practices for waived testing
Notes: Tip: Add your own speaker notes here.

## Slide 4: What is a “waived” test?
Under CLIA, tests are categorized by the difficulty to perform the test
Waived tests are:
  Simple lab examinations or procedures
  Cleared by FDA for home use
  Negligible likelihood of erroneous results
  No reasonable risk of harm if performed incorrectly
Notes: Tip: Add your own speaker notes here.

## Slide 5: What does “waived” mean?
Waived from most of CLIA requirements established
for moderate and high complexity
  Site inspections are not routinely performed
  Proficiency testing is not required
  CLIA personnel qualifications do not apply
Notes: Tip: Add your own speaker notes here.

## Slide 6: Pilot Studies on Waived Testing
1999-2001 (n=500)
50% of the labs do not follow manufacturer’s instructions or do not have product insert (PI)
Other findings indicated that these labs had test performance-related problems
Section 5- results of the pilot study

## Slide 7: Recommendations of the Pilot study
Institute an educational program
Validate the effectiveness of this educational program
Survey a percentage of waived laboratories annually

## Slide 8: Certificate of Waiver (COW) Surveys
April 2002 CMS initiated on-site visits to approx. 2% of COW labs in each State
Random selection
Use of a questionnaire focusing on laboratory practices
Educational
CDC maintains the data base of the results obtained from these surveys

## Slide 9: CMS COW Study-results
CMS surveys 2002-2004 (n=4214)
12% did not have current instructions/PIs
21% did not routinely check for changes in the instructions
RE: following the manufacturer’s instructions
  sites did not
  perform quality control 21%
  adhere to expiration dates 6%
  use appropriate specimen 2%

## Slide 10: How to find if your kit or test system is waived
CLIA Website
http://www.cms.gov/clia/
FDA Websites
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfClia/
analyteswaived.cfm
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfCLIA/search.cfm

## Slide 11: CLIA Requirements for Waived Testing
Enroll in the CLIA program
Follow manufacturer’s test instructions
Pay applicable certificate fees biennially
By the CLIA law, COW labs perform only tests that are determined by FDA or CDC to be so simple that there is little risk of error

## Slide 12: Manufacturer’s instructions or Package/Product Insert
You must have a current PI for waived tests
You must read the instructions in the PI
You must follow instructions in the PI explicitly
You must evaluate results and take appropriate action

## Slide 13: Test procedure/Directions/Instructions
Follow exactly
Don’t modify
Adhere to timing

## Slide 14: PI: Focus on these sections
Intended use Procedural notes
Product storage Quality control
Precautions/ Results/
warnings interpretation
Patient preparation Limitations
Specimen collection Expected results
Test procedure

## Slide 15: Regulatory Terminology
“shall”
“must”
“test”
“do”
“perform”
“follow”
“require”
“always”

## Slide 16: Terms not regulatory
“should”
“recommend”

## Slide 17: Good Laboratory Practices (GLPs)
Important to understand the concept of GLP-Why?
Assure correct/quality test results
List of GLPs included in binder
Can you think of some?

## Slide 18: Good Laboratory Practices for Waived Testing
Document the name of the test, lot number, and expiration date
Routinely review incoming package inserts for changes and updates by the manufacturer
Testing personnel training
Notes: Any laboratory cannot send it PT samples out to another lab for analysis.
PT referral is referred to as “IMPROPER REFERRAL” in the CLIA Law.
“Intentional” mean in some court ruling: “Intentionally referred” [as in “intentionally referred” its proficiency testing samples to another laboratory for analysis] requires not specific intent, but general intent, that is an intent to act. Motive is irrelevant.
Strict penalties: automatic REVOCATION of CLIA certificate, lab cannot test for one year, LD prohibited to direct lab for two years.

## Slide 19: Good Laboratory Practices for Waived Testing-con’t.
Policy/procedure for checking patient ID and documenting result in patient’s chart
Run quality controls, participate in proficiency testing
Instrument or device error codes
Initials of person who performed the tests

## Slide 20: Good Laboratory Practices for Waived Testing-con’t.
What are built-in, internal procedural controls?
Example of Strep A Rapid Test Device (Section 3 in your binder)
  Internal Positive Quality Control confirms
  1. sufficient specimen volume
  2. correct procedural technique
  Internal Negative Quality Control confirms the test is working properly

## Slide 21: Good Laboratory Practices for Waived Testing-con’t.
What are external controls?
(Previous example)
External positive/negative controls:
1. May be provided with the kit/test
2. Will check the entire kit/test system
3. Confirm your kit/test is working and is accurate
4. Should be run once per kit for good lab practice

## Slide 22: THANK YOU!
Contacts:
  Doris Thompson doris.thompson@alaska.gov
  907-334-2583
  Fran Lehr francisca.lehr@cms.hhs.gov
  206-615-2710