## Slide 1: STREAMLINED QC
Debra Kieler, BS, MT(ASCP)
May 10, 2012

## Slide 2
Southwest Health Center (SHC)-Platteville, WI
25 bed critical access hospital, 84 bed nursing home, and 10 bed inpatient adult psychiatric unit
Hospital lab staffed 24/7/366 by MLT, MT, and MLS personnel
Microbiology mainly 0600 to 1800 daily
ID and susceptibilities of aerobic gram negative and gram positive organisms
Biomerieux Vitek 2 Compact-purchased in 2006
Gram negative ID and AST cards
Gram positive ID and AST cards
ID and AST reporting with Vitek 2 - 2007

## Slide 3
AST/MIC QC-each new lot of cards/shipment and weekly thereafter
AST-GP67-gram positive organisms
  Staph aureus-ATCC 29213
  Staph aureus-ATCC BAA1026
  Staph aureus-ATCC BAA 977
  Staph aureus-ATCC BAA 976
  Enterococcus faecalis-ATCC 51299
  Enterococcus faecalis-ATCC 29212
AST-GN45-gram negative organisms
  E. coli ATCC 25922
  E. coli ATCC 35218
  Pseudomonas aeruginosa ATCC 27853
  Klebsiella pneumoniae ATCC 700603

## Slide 4
This is the minimum suggested QC for susceptibility testing
NO streamlining AST QC on Vitek 2 Compact
But, streamline qc in micro with ID cards for gram negative and gram positive organisms or GN and GP cards, is possible

## Slide 5
CLSI M50-A Volume 28, No. 23- Quality Control for Commercial Microbial Identification Systems; Approved Guideline document
Biomerieux (BMX) communications

## Slide 6
MIS-Microbial Identification Systems
MIS-test systems that utilize multiple substrates and/or reagents to identify aerobic or anaerobic bacteria, yeast, moulds, or yeast-like algae grown from culture
CLIA regulations effective 9/1/1992-QC of a commercial MIS-every reagent/and/or substrate of each batch, lot number, and shipment when prepared or opened for positive and negative reactivity (42CFR 493.1256); interpretive guidelines (42CFR 493.1261) also state that the lab must use control organisms to verify positive and negative reactivity.
Ongoing QC testing validates that the MIS continues to perform acceptably throughout its use

## Slide 7
Satisfy CLIA for MIS-extensive QC testing even though the frequency of failures is low; retesting that is of questionable value
2005-ASM survey showed real product failure was extremely rare; test reactions of specific substrates rather than all substrates; use key indicator strains; detect degradation of labile reagents and/or substrates
Not all manufacturers provide this-these systems must follow comprehensive CLIA QC requirements

## Slide 8
STREAMLINED QC-the reduction of QC for a commercial MIS as compared to the requirements in the CLIA regulations
NOTE: The manufacturer of the MIS will specify the key indicator strains and reagents/substrates to be tested; provide the expected QC results based on historical data and ongoing QA programs at the site of manufacture. The lab user MAY ADD but MAY NOT DECREASE the number of specific QC organisms and/or reagents/substrates to be tested and reviewed as recommended by the manufacturer for each lot and shipment. Implementation of Streamlined QC by the user can only occur after the user has verified the performance of the test system in use.

## Slide 9
Key roles-Manufacturer, Distributor, User
MANUFACTURER
Follow quality system standards and regulations
Provide adequate instructions for use
Selection of QC organisms appropriate for their MISs
Acceptance criteria for QC testing
Explain why strains were selected
Provide statement of compliance with ISO 13485 and QSRs to the user
Maintain complaint files
Investigate, document, and monitor complaints to determine if the MIS still meets performance characteristics

## Slide 10
DISTRIBUTOR
Comply with manufacturer’s specs for MIS storage and handling before and during delivery to the user
Maintain documentation that verifies compliance-available to manufacturer and user if requested
Communicate user complaints to the manufacturer
USER
Storage and handling on receipt and before testing is followed
Staff –qualified, trained, and competent to use MIS
Retain/follow current technical information/product insert
Document all MIS QC activities and corrective action
Comply with all regulations for testing and retaining documents-local, state, and federal requirements

## Slide 11: Criteria for Initiation of Streamlined QC
      Verification of performance-CLIA-42 CFR 493 1253(b)-been conducted?
      Performing comprehensive CLIA QC of each reagent/substrate?
      Under CLIA, verification study required after 9/1/92-documentation of study conducted
      Before 9/1/92-if verification study not done or no documentation of verification after this date but has been doing comprehensive CLIA QC since 1992 on an ongoing basis, historical review of its QC meets qualification

## Slide 12
If a lab has not conducted a verification study and has not been performing comprehensive CLIA QC, it does not qualify for streamlined QC until it has documentation that it has met one of these criteria.
A commercial MIS that is modified by the user does not qualify for streamlined QC either.

## Slide 13: INITIAL QUALIFICATION
1) Maintain current documentation of manufacturer’s conformance with ISO 13485 and QSR requirements
2) Meet ONE of the following:
a) If verification study performed and performance specifications comparable to the manufacturer-
STREAMLINED QC MAY BE DONE IMMEDIATELY
(Documentation of the verification study must be available)
b) If verification study not performed or lacking required documentation but comprehensive QC has been performed, user may conduct and document a historical review of QC performance
as follows:

## Slide 14
Review QC performance for at least 3 consecutive lot numbers of the MIS from 3 different shipments that span at least 3 consecutive seasons to assess seasonal variation of shipping conditions.
QC testing must be performed using positive and negative controls for each reagent and/or substrate according to the manufacturer’s instructions.
Performance shall be considered satisfactory and the user may initiate streamlined QC if 95% of the reagent/substrate results are within the results specified by the manufacturer. If the QC meets the 95% threshold, but the same reagent/substrate fails repeatedly, it may indicate a problem. These results should be further evaluated.
If the user has insufficient data or unexpected QC results occur, consult CLIA regs before proceeding.

## Slide 15: STREAMLINING GP CARD QC AT SHC
GP CARD QC initially (2006 to 2011)
Enterococcus casseliflavus ATCC 700327
Kocuria kristianae ATCC BAA -752
Listeria monocytogenes ATCC BAA -751
Staph aureus spp.aureus ATCC 29213
Staph saprophyticus ATCC BAA -750
Staph scuiri ATCC 29061
Strep equi ssp.zooepidemicus ATCC 43079
Strep thermophilus ATCC 19258
Shipment 1- 10/2006 to 08/2007
Shipment 2- 08/2007 to 01/2009
Shipment 3-01/2009 to 07/2010
Shipment 4-07/2010 to 12/2011

## Slide 16
All expected positive, negative, or variable results had to be obtained
Any out-of- control results had to be repeated until satisfactory
43 biochemicals tested per qc organism
#1, #2, and #3 shipments performed as expected
#4 shipment had 0129R (-) vs. (+) on 19258 and AspA (-) vs. (+) on 700327
Repeat testing produced (+) results as expected
July, 2009 - BMX customer letter informed us of organisms selected for new streamlined qc and new comprehensive qc requirements so we could modify our system per new CLSI M50-A.
Enterococcus casseliflavus ATCC 700327
Strep thermophilus ATCC 19258
These QC organisms were already being used previously so in stock here.

## Slide 17
We were ready to start in Dec., 2011
Documentation that BMX complies with ISO 13485 and QSR-yes
Verification study-no
Comprehensive QC –yes (Three consecutive lot numbers; three different shipments; three consecutive seasons)
QC testing using positive and negative on each reagent/substrate-yes
95% of the reagent/substrate with results specified by BMX-yes
E. casseliflavus (700327) and Strep thermophilus (19258)-yes
All reactions as expected-yes
INITIAL QUALIFICATION HAS BEEN MET-12/2011

## Slide 18
BMX changed QC organisms requirements- 11 or 12/2011
Staph saprophyticus (ATCC BAA-750 ) replaced Strep thermophilus (ATCC 19258).
CONTINUED QUALIFICATION-12/2011 and beyond
Maintain current documentation of manufacturer’s conformance to QSR requirements
Maintain documentation of verification study or historical QC review
Test all key indicator strains with each batch, lot number, and shipment
Perform testing according to manufacturer’s instructions
Monitor and document streamlined QC performance
Investigate and resolve any QC failures
Report QC failures to the manufacturer and distributor

## Slide 19: STREAMLINING GN CARD QC AT SHC
GN CARD QC initially (2006-2010)
Acinetobacter baumannii ATCC BAA-747
Enterobacter cloacae ATCC 700323
Klebsiella oxytoca ATCC 700324
Ochrobactrum anthropi ATCC BAA-749
Proteus vulgaris ATCC 6380
Shigella sonnei ATCC 25931
Stenotrophomonas maltophilia ATCC 17666
Shipment 1-10/2006 to 05/2007 Shipment 5- 08/2008 to 03/2009
Shipment 2- 05/2007 to 08/2007 Shipment 6- 03/2009 to 05/2009
Shipment 3- 08/2007 to 03/2008 Shipment 7- 05/2009 to 10/2009
Shipment 4- 03/2008 to 08/2008 Shipment 8- 10/2009 to 04/2010

## Slide 20
All expected +, -, or v results had to be obtained
Any out-of-control results had to be repeated until correct
47 biochemicals tested per organism
Shipment 1-1 failure; repeat ok
Shipment 2-8 failures; repeats ok
Shipment 3-6 failures; repeats ok
Shipment 4-1 failure; repeat ok
Shipment 5-2 failures; repeats ok
Shipment 6- 1 failure; repeat ok
Shipment 7- all ok
Shipment 8- 1 failure; repeat ok

## Slide 21
Shipment 9- 04/2010
Ochrobactrum anthropi - BAA-749 proved to be a challenge even if it was April Fool’s Day. TyrA and AGLU failed repeatedly; we even enlisted BMX’s help and ended up substituting an alternate organism to obtain the results for TyrA and AGLU to have the correct reactions expected in comprehensive qc. In 08/2010-BMX changed the expected value from + to v for AGLU.
WE WERE EXTREMELY MOTIVATED TO GO TO STREAMLINING QC ON SHIPMENT 10!

## Slide 22
Shipment 10- 10/2010
July, 2009-BMX customer letter informed us of organisms selected for new streamlined qc and new comprehensive qc requirements so we could modify our system per new CLSI M50-A.
Enterobacter hormaechei ATCC 700323
(formerly Enterobacter cloacae)
Stenotrophomonas maltophilia ATCC 17666
These QC organisms were already being used previously so in stock here.

## Slide 23
We were ready to start in Oct., 2010
Documentation that BMX complies with ISO 13485 and QSR-yes
Verification study-no
Comprehensive qc-yes (Three consecutive lot numbers; three different shipments; three consecutive seasons)
QC testing using positive and negative on each reagent/substrate-yes
95% of the reagent/substrate with results specified by BMX-yes
Enterobacter hormaechei (700323) and Stenotrophomonas maltophilia (17666)-yes
All reactions as expected-yes
INITIAL QUALIFICATION HAS BEEN MET-10/2010

## Slide 24
CONTINUED QUALIFICATION-10/2010 and beyond
Maintain current documentation of manufacturer’s conformance to QSR requirements-yes
Maintain documentation of verification study or historical QC review-yes
Test all key indicator strains with each batch, lot number, and shipment-yes
Perform testing according to manufacturer’s instructions-yes
Monitor and document streamlined QC performance-yes
Investigate and resolve any QC failures
Report QC failures to the manufacturer and distributor

## Slide 25
Streamlined QC -10/10 to now
Shipments received 10/10, 07/11, and 02/12
All shipments performed as expected with no failures
Streamlined QC has proven to save time and money at SHC!

## Slide 26: Questions?
Contact Debra Kieler
Email at andrewsg@southwesthealth.org
Phone at (608) 342-4760
Fax at (608) 342-5094